## Cal Data Report Template
Tektronix Inc.
13725 SW Karl Braun Drive
Beaverton, OR 97077
503-627-7111
CALIBRATION DATA REPORT
|  | Certificate Number: | 1123 |  |  |  | Revision: | 0
|  | Manufacturer: | Model: |  | Serial Number: |  | Calibration Date:
|  | Tektronix | DPO1202 |  | PQCNUD001123 |  | 19/03/2026
The data in this report was recorded at an automated calibration station.
The data may differ in step content from that produced by completing the Performance Verification tests in the Service Manual or the User Manual.
Either set of data verifies instrument conformance to Advertised Specifications.
CALIBRATION DATA REPORT
|  | Certificate Number: | 1123 |  |  |  | Revision: | 0
SPECIFICATION TESTED: Signal Path Gain
|  |  | Voltage | Low | Measured | High
|  | Channel # | Scale | Limit | Value | Limit | Units | Result
|  | CH1 | 1 mV/div | 5.88 |  | 6.12 | mV | Pass
|  |  | 2 mV/div | 11.76 |  | 12.24 | mV | Pass
|  |  | 5 mV/div | 29.4 |  | 30.6 | mV | Pass
|  |  | 10 mV/div | 59.1 |  | 60.9 | mV | Pass
|  |  | 20 mV/div | 118.2 |  | 121.8 | mV | Pass
|  |  | 50 mV/div | 295.5 |  | 304.5 | mV | Pass
|  |  | 100 mV/div | 591 |  | 609 | mV | Pass
|  |  | 200 mV/div | 1.182 |  | 1.218 | V | Pass
|  |  | 500 mV/div | 2.955 |  | 3.045 | V | Pass
|  |  | 1 V/div | 5.91 |  | 6.09 | V | Pass
|  |  | 2 V/div | 11.82 |  | 12.18 | V | Pass
|  |  | 5 V/div | 29.55 |  | 30.45 | V | Pass
|  |  | 10 V/div | 59.1 |  | 60.9 | V | Pass
|  | CH2 | 1 mV/div | 5.88 |  | 6.12 | mV | Pass
|  |  | 2 mV/div | 11.76 |  | 12.24 | mV | Pass
|  |  | 5 mV/div | 29.4 |  | 30.6 | mV | Pass
|  |  | 10 mV/div | 59.1 |  | 60.9 | mV | Pass
|  |  | 20 mV/div | 118.2 |  | 121.8 | mV | Pass
|  |  | 50 mV/div | 295.5 |  | 304.5 | mV | Pass
|  |  | 100 mV/div | 591 |  | 609 | mV | Pass
|  |  | 200 mV/div | 1.182 |  | 1.218 | V | Pass
|  |  | 500 mV/div | 2.955 |  | 3.045 | V | Pass
|  |  | 1 V/div | 5.91 |  | 6.09 | V | Pass
|  |  | 2 V/div | 11.82 |  | 12.18 | V | Pass
|  |  | 5 V/div | 29.55 |  | 30.45 | V | Pass
|  |  | 10 V/div | 59.1 |  | 60.9 | V | Pass
CALIBRATION DATA REPORT
| Certificate Number: | 1123 | [7 empty] | Revision: | 0
SPECIFICATION TESTED: Bandwidth
|  | Voltage | Low | Measured Value |  |  |  |  | High
| Channel # | Scale | Limit | 10MHz | 20MHz | 100MHz | 150MHz | 200MHz | Limit | Units | Result
| CH1 | 1 mV/div | 4.242 | [5 empty] | 6.36 | mV | Pass
|  | 2 mV/div | 8.484 | [5 empty] | 12.72 | mV | Pass
|  | 5 mV/div | 21.21 | [5 empty] | 31.8 | mV | Pass
|  | 10 mV/div | 42.42 | [5 empty] | 63.6 | mV | Pass
|  | 20 mV/div | 84.84 | [5 empty] | 127.2 | mV | Pass
|  | 50 mV/div | 212.1 | [5 empty] | 318 | mV | Pass
|  | 100 mV/div | 424.2 | [5 empty] | 636 | mV | Pass
|  | 200 mV/div | 848.4 | [5 empty] | 1272 | mV | Pass
|  | 500 mV/div | 2.121 | [5 empty] | 3.18 | V | Pass
|  | 1 V/div | 4.242 | [5 empty] | 6.36 | V | Pass
|  | 2 V/div | 4.242 | [5 empty] | 6.36 | V | Pass
|  | 5 V/div | 4.242 | [5 empty] | 6.36 | V | Pass
| CH2 | 1 mV/div | 4.242 | [5 empty] | 6.36 | mV | Pass
|  | 2 mV/div | 8.484 | [5 empty] | 12.72 | mV | Pass
|  | 5 mV/div | 21.21 | [5 empty] | 31.8 | mV | Pass
|  | 10 mV/div | 42.42 | [5 empty] | 63.6 | mV | Pass
|  | 20 mV/div | 84.84 | [5 empty] | 127.2 | mV | Pass
|  | 50 mV/div | 212.1 | [5 empty] | 318 | mV | Pass
|  | 100 mV/div | 424.2 | [5 empty] | 636 | mV | Pass
|  | 200 mV/div | 848.4 | [5 empty] | 1272 | mV | Pass
|  | 500 mV/div | 2.121 | [5 empty] | 3.18 | V | Pass
|  | 1 V/div | 4.242 | [5 empty] | 6.36 | V | Pass
|  | 2 V/div | 4.242 | [5 empty] | 6.36 | V | Pass
|  | 5 V/div | 4.242 | [5 empty] | 6.36 | V | Pass
CALIBRATION DATA REPORT
|  | Certificate Number: | 1123 |  |  |  | Revision: | 0
SPECIFICATION TESTED: AFG Output Offset
|  |  | Low | Measured | High
|  | Target DC Voltage | Limit | Value | Limit | Units | Result
|  | -2.5 | -2.627 | N/A | -2.373 | V | Pass
|  | -2 | -2.102 | N/A | -1.898 | V | Pass
|  | -1.5 | -1.577 | N/A | -1.423 | V | Pass
|  | -1 | -1.052 | N/A | -0.948 | V | Pass
|  | -0.5 | -0.527 | N/A | -0.473 | V | Pass
|  | 0 | -0.002 | N/A | 0.002 | V | Pass
|  | 0.5 | 0.473 | N/A | 0.527 | V | Pass
|  | 1 | 0.948 | N/A | 1.052 | V | Pass
|  | 1.5 | 1.423 | N/A | 1.577 | V | Pass
|  | 2 | 1.898 | N/A | 2.102 | V | Pass
|  | 2.5 | 2.373 | N/A | 2.627 | V | Pass
SPECIFICATION TESTED: Time base accuracy
|  |  |  | Low | Measured | High
|  | Test Condition |  | Limit | Value | Limit | Units | Result
|  | Time base gear/div | 5 ns/div | -12.5 |  | 12.5 | ns | Pass
|  | time marker frequency | 1 kHz
|  | Period of signal | 1 ms
|  | Time base delay | 1 ms

## After Sales Template (hidden)
Tektronix Inc.
13725 SW Karl Braun Drive
Beaverton, OR 97077
503-627-7111
CALIBRATION DATA REPORT
| Report Type: | <indicate the report type here>
| Certificate Number: | <certificate number> |  |  |  |  | Revision: | 0
| Manufacturer: | Model: |  |  | Serial Number: |  | Calibration Date:
| Tektronix | <model number> |  |  | <xxxxxx> |  | <date of calibration>
The data in this report was recorded at an automated calibration station.
The data may differ in step content from that produced by completing the Performance Verification tests in the Service Manual or the User Manual.
Either set of data verifies instrument conformance to Advertised Specifications.
CALIBRATION DATA REPORT
| Certificate Number: |  | <certificate #> |  |  |  | Revision: | 0
| SPECIFICATION TESTED: |  |  | <xxxxxxxxxxxxx>
|  |  |  | Low | Measured | High
| Step/Operation |  |  | Limit | Value | Limit | Units | Result
| <xxxxxxxxxx> |  |  | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <Pass/Fail>
| <xxxxxxxxxx> |  |  | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <Pass/Fail>
| <xxxxxxxxxx> |  |  | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <Pass/Fail>
| <xxxxxxxxxx> |  |  | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <Pass/Fail>
| <xxxxxxxxxx> |  |  | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <Pass/Fail>
| <xxxxxxxxxx> |  |  | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <Pass/Fail>
| <xxxxxxxxxx> |  |  | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <Pass/Fail>
| <xxxxxxxxxx> |  |  | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <Pass/Fail>
| <xxxxxxxxxx> |  |  | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <Pass/Fail>
| <xxxxxxxxxx> |  |  | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <xxxxxxx> | <Pass/Fail>

## Revision Control (hidden)
REVISION RECORD
Version | Issue Date | Initiator | Nature of Change
1 | 2012-07-23 | D. Knapp | Initial Release
2 | 2013-08-15 | S. Moskalets | Minor format update