## Sheet1
Room Temperature Performance
Ambient temperature: 23℃±1℃ |  |  |  |  | Product version
Ambient humidity: <80%RH
Test location:
Equipment used: | FLUKE9500B oscilloscope calibration instrument, UTG9000T signal generator, MSO3000HD oscilloscope<afgEqu>.
Tester: |  |  |  |  | Test date:
Test item 1: Bandwidth Test
Test requirements: Input standard reference sine wave frequency of 1 MHz and Bandwidth Frequencies from table at 6 divisions peak-to-peak amplitude from a standard signal source to each of the channels listed in the table below with a 50 ohm through termination. Adjust the oscilloscope's timebase and trigger level to display approximately 5 cycles of the waveform on the screen (at stable triggered). Record the reference frequency amplitude for comparison to subsequent Bandwidth Frequencies. Record the amplitude of each frequency point in half-channel, and full-channel modes respectively, and record the test result in the following table. At V/div 1V and above, set generator to max setting and proceed with measurement (Measurement will only consume a few vertical Divisions).
Index requirements: bandwidth ≥-3dB/0.707 of the measurement value;
Single Channel Test |  |  |  | unmeasured
Vertical gear |  |  |  | 1mV | 2mV | 5mV | 10mV | 20mV | 50mV | 100mV | 200mV | 500mV | 1V | 2V | 5V
Input value(Vin_pp) |  |  |  | 6 | 12 | 30 | 60 | 120 | 300 | 600 | 1.2 | 3.00 V | 5.50 V | 5.50 V | 5.50 V
<model> | CH1 | 10MHz Measured value Vbw_pp
Gain
20MHz Measured value Vbw_pp
Gain
50MHz Measured value Vbw_pp
Gain
|  | Result |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
| CH2 | 10MHz Measured value Vbw_pp
Gain
20MHz Measured value Vbw_pp
Gain
50MHz Measured value Vbw_pp
Gain
[6 empty] | 0.8713 |  |  | 0.8675 |  | 0.825
|  | Result |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
| CH3 | 10MHz Measured value Vbw_pp
Gain
20MHz Measured value Vbw_pp
Gain
50MHz Measured value Vbw_pp
Gain
|  | Result |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
| CH4 | 10MHz Measured value Vbw_pp
Gain
20MHz Measured value Vbw_pp
Gain
50MHz Measured value Vbw_pp
Gain
|  | Result |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
Dual-channel mode: one channel has input signal when both channels are turn on, for example: turn on CH1 and CH2, CH1 has input signal, CH2 do not input signal
Vertical gear |  |  |  | 1mV | 2mV | 5mV | 10mV | 20mV | 50mV | 100mV | 200mV | 500mV | 1V | 2V | 5V
Input value(Vin_pp) |  |  |  | 6 | 12 | 30 | 60 | 120 | 300 | 600 | 1.2 | 3.00 V | 5.50 V | 5.50 V | 5.50 V
<model> | CH1 | 10MHz Measured value Vbw_pp
Gain
20MHz Measured value Vbw_pp
Gain
50MHz Measured value Vbw_pp
Gain
|  | Result |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
| CH2 | 10MHz Measured value Vbw_pp
Gain
20MHz Measured value Vbw_pp
Gain
50MHz Measured value Vbw_pp
Gain
|  | Result |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
| CH3 | 10MHz Measured value Vbw_pp
Gain
20MHz Measured value Vbw_pp
Gain
50MHz Measured value Vbw_pp
Gain
|  | Result |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
| CH4 | 10MHz Measured value Vbw_pp
Gain
20MHz Measured value Vbw_pp
Gain
50MHz Measured value Vbw_pp
Gain
|  | Result |  | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
Test item 2: DC Gain accuracy
Test requirements: Center the waveform base line, Set the Fluke 9500B/C calibrator output to DC voltage at +3div, measure the waveform average under average acquisation mode with 16 times average number. Then sbstract the measurement value with that of the -3div DC input and record the result for every vertical scale in the following table.
Index requirement: <5mV/div: ±3.0%, derated at 0.100%/°C above 30°C. ≥5mV/div: ±2%, derated at 0.100%/°C above 30°C.
Vertical gear |  |  |  | 1mV | 2mV | 5mV | 10mV | 20mV | 50mV | 100mV | 200mV | 500mV | 1V | 2V | 5V | 10V
Input value(mV/VDC) |  |  |  | 6 | 12 | 30 | 60 | 120 | 300 | 600 | 1.2 | 3 | 6 | 12 | 30 | 60
Error range(±) |  |  |  | 3% | 3% | 2% | 2% | 2% | 2% | 2% | 2% | 2% | 2% | 2% | 2% | 2%
<model> (Single Channel On) | CH1 |  | +DC
-DC
DC difference
Measurement error
|  |  | Result | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
| CH2 |  | +DC
-DC
DC difference
Measurement error
|  |  | Result | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
| CH3 |  | +DC
-DC
DC difference
Measurement error
|  |  | Result | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
| CH4 |  | +DC
-DC
DC difference
Measurement error
|  |  | Result | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
Test item 3: Time base accuracy test
Test requirement: 1、Input Configuration: Set the Fluke 9500B/C channel 1 (CH1) to a 1 kHz, 3 Vpp time marker. Adjust the timebase to the values listed in the table. 2、Adjust the trigger level to the center of the waveform. 3、Measurement: Measure the time difference (△T) between the waveform edge and the horizontal center of the screen.
Technology index：25 ppm (from the latest PRS)
Machine type |  | Time base gear/div | 5ns/div
|  | input square wave frequency | 1kHz
|  | Period of signal | 1ms
|  | Time base delay | 1ms
|  | Technology index： | 25
<model> |  | Measured value/ps
|  | Result | PASS
Test item 4: DC Offset Accuracy (For -S options only):
Test procedure: 1.Connect Output of Function generator to input channel of DMM 2.Set the function generator to DC, the output impedance to the high-impedance state, and the DC voltage set to below table. 3.The digital multimeter sets the DCV measurement function, measurement result are recored in below table. | [16 empty] | Low oversampling rate, for reference only
Index requirements: ± (offset set value 5%+2 mV)
Nominal value（V） | -2.5 | -2 | -1.5 | -1 | -0.5 | 0 | 0.5 | 1 | 1.5 | 2 | 2.5
index（V） | 0.123 | 0.098 | 0.073 | 0.048 | 0.023 | 0.002 | 0.027 | 0.052 | 0.077 | 0.102 | 0.127
Measured DC voltage（V） | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Test results | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS | PASS
Test item 5: input impedance test
Test requirements: The input impedance is measured separately in the corresponding gear （Reading from FLUKE 9500B)
Technical index: (1MΩ± 2%) ¦¦ (17 pF+1/-2 pF)
Vertical gear |  |  | 50 |  |  |  | 500 |  |  |  | 5V
|  |  | input resistance(1MΩ) |  | input capacitance(pF) |  | input resistance(1MΩ) |  | input capacitance(pF) |  | input resistance(1MΩ) |  | input capacitance(pF)
<model> (Single Channel On) | CH1 | Measured value
| CH2 | Measured value
| CH3 | Measured value
| CH4 | Measured value
| Measuring result |  | PASS |  | PASS |  | PASS |  | PASS |  | PASS |  | PASS
unmeasured
Test item 6: DC offset test
Test requirements: Set the generator output to DC voltage (FLUKE 9500B), the signal generator set +DC voltage, the oscilloscope offset will need to set to -DC voltage; for example: the oscilloscope offset is +2V, signal generator set to -2V. Measure at the separation value of the baseline
Range of bias voltage for example： 1mV/div to 50 mV/div：±2 V 51 mV/div to 1 V/div：±25 V 1.02 V/div to 10 V/div：±250 V
Offset accuracy: ±0.1div ±2mV ±2% offset
Vertical gear(mV) |  |  | 50 |  | 500 |  | 5
Voltage of offset(unit：mV） |  |  | 2000 mV | -2000 | 25000 | -25000 | 60000 | -60000
<model> (Single Channel On) | CH1 | Measured value(mV)
|  | Measuring result | PASS | PASS | PASS | PASS | PASS | PASS
| CH2 | Measured value(mV)
|  | Measuring result | PASS | PASS | PASS | PASS | PASS | PASS
| CH3 | Measured value(mV)
|  | Measuring result | PASS | PASS | PASS | PASS | PASS | PASS
| CH4 | Measured value(mV)
|  | Measuring result | PASS | PASS | PASS | PASS | PASS | PASS
Test item 7: Baseline noise test
Test requirements: Select the measured channel, connect 50Ω impedance match at the input, turn off all the other channels; set the channel to DC coupling, observe the baseline noise at the highest sampling rate, set to auto memory depth ; and maximum waveform brightness.
Input Impedance 1MΩ; (Measurement Result is for reference only)
Machine type | channel | Vertical gear/div | 1mV |  | 2mV |  | 100mV
|  | Technology index (AC Vrms) | 200 |  | 200 |  | 4000
<model> (Single Channel On) | CH1 | Measured value(50ns/div)
Measured value(1ms/div)
Measured value(10ms/div)
|  | unit (ACVrms) | μVrms |  | μVrms |  | mVrms
actual width /div
| CH2 | Measured value(50ns/div)
Measured value(1ms/div)
Measured value(10ms/div)
|  | unit (ACVrms) | μVrms |  | μVrms |  | mVrms
actual width /div
| CH3 | Measured value(50ns/div)
Measured value(1ms/div)
Measured value(10ms/div)
|  | unit (ACVrms) | μVrms |  | μVrms |  | mVrms
actual width /div
| CH4 | Measured value(50ns/div)
Measured value(1ms/div)
Measured value(10ms/div)
|  | unit (ACVrms) | μVrms |  | μVrms |  | mVrms
actual width /div
Test item 8: channel latency test
Test requirements: 1. Select two measurement channels (For example CH1 and CH2), and the Settings should be consistent, the electrical characteristics of the connecting cables of the two channels should be consistent, and close the other channels (for example CH3 and CH4) ; 2. Set the input signale to 100MHz , 300mVpp; and set CH1 and CH2 to be AC coupling, timebase to be 2ns/div, 50mV/div vertical scale; 3. Use cursor time measurement to measure the time difference between adjacent edges of two channels and record it as t1; 4. Exchange the two input signals, measure the time difference between the adjacent edges of the two channels again, and record it as t2; The delay time △t, △t=(t1+t2)/2
Technology index：Delay Delta Time <=500ps
Machine type |  | channel | Group 1 |  | Group 2
|  |  | CH1 | CH2 | CH3 | CH4
<model> |  | Technology index (ps) | 500 |  | 500
Measured value (ps)
|  | Result | PASS |  | PASS
Test item 9:Sinewave test (-S AFG Option Only)
Test requirements: Measure using Frequency Counter. Measure output set to 10kHz Sine, full amplitude.
Technology index：130ppm (frequency ≤10 KHz); 50ppm (frequency > 10 KHz)
Machine type | Waveform |  | 10kHz Sine | 25MHz Sine
<model> (Single Channel On) | Technology index |  | +/- 1Hz | +/- 1kHz
| Measured value (Hz) |  | N/A | N/A
| Result |  | PASS | PASS
Test item 10: Squarewave test (-S AFG Option Only)
Test requirements: Measure using Frequency Counter. Measure output set to 10kHz Sine, full amplitude.
Technology index：130ppm (frequency ≤10 KHz); 50ppm (frequency > 10 KHz)
Machine type | Waveform |  | 25kHz Square | 10MHz Square
<model> (Single Channel On) | Technology index |  | +/- 4Hz | +/- 400Hz
| Measured value (Hz) |  | N/A | N/A
| Result |  | PASS | PASS
Test Item 11: USB/LAN Connection Reliability Test
Test requirements: Send multiple SCPI commands from the computer to the machine (not system setting) Observe that the machine under test does not causing system crash or other abnormal behavior, and ensure that communication does not interrupt.
Prototype number | Interface type | Test results | Remarks
Machine No. 1 | USB | PASS
Machine No. 2 | LAN | PASS
Test Item 12: AUXOUT Output Voltage Test
Test requirements: Measure output voltage on AUXOUT Connector Characteristic Limits Vout (HI) ≥ 2.5 V open circuit; ≥ 1.0 V into a 50Ω load to ground. Vout (LO) ≤ 0.7 V into a load of ≤ 4 mA; ≤ 0.25 V into a 50Ω load to ground.
open circuit | [5 empty] | 50Ω
Machine type | AUXOUT SETTING |  | HIGH | LOW |  | Machine type | AUXOUT SETTING |  | HIGH | LOW
<model> | Measured value (V) |  |  |  |  | <model> | Measured value (V)
| Result |  | PASS | PASS |  |  | Result |  | PASS | PASS
Test Item 13: Probe Compensator Output Voltage Test
Test requirements: Measure output voltage on Probe Compensator Characteristic Output Voltage: 2.5V +/-5%, Frequency 1kHz +/- 25%
Machine type | Probe Comp |  | Voltage | Frequency
<model> | Measured value (V, kHz)
| Result |  | PASS | PASS
Test Item 14: Screen functionality test
Test requirements: Test functionality of RGB and backlight of screen, as well as touch panel.
Machine type | Screen Test |  | Display | Touch
<model> | Result |  | PASS | PASS